Introduction to pH

Standardization of the pH meter

(This is a general procedure.  It is preferable to use the manufacturer's procedure for standardization.)

1.	Set the temperature control to the temperature of the laboratory.

2.	Insert the electrode into the pH 7 buffer, and adjust the pH to 7.00 with the standardization knob.  Turn the function knob to stand by.  Remove the electrode from the solution and place over a waste beaker.  Rinse with deionized water and wipe with a Kimwipe or a paper towel.  If the solution to be measured is acidic, continue with step 3.  If the solution is basic, proceed to step 4.

3.	Insert the electrode into the pH 4 buffer, turn the function knob to pH,  and adjust the pH to 4.00 with the slope knob.  Turn the function knob to stand by.  Remove the electrode from the solution and place over a waste beaker.  Rinse with deionized water and wipe with a Kimwipe or a paper towel.  Proceed to step 5.

4.	Insert the electrode into the pH 10 buffer and adjust the pH to 10.00 with the slope knob. Turn the function knob to stand by.  Remove the electrode from the solution and place over a waste beaker.  Rinse with deionized water and wipe with a Kimwipe or a paper towel.

5.	Repeat until the best range is achieved.

Last updated 8-01